## hard ref
称重传感器 | 中航电测 | L6D
| 等级 | C3
| 灵敏度 | 2.0+-0.2 mV/V
| 量程 | 50Kg
| 极限过载 | 150%
| 激励电压范围 | 10VDC - 15VDC | 典型值10V
称重传感器数据处理 | 20MV放大到4V | 倍率 | 200倍 4V
| 仪表放大器 | 单电源AD623
|  | AD8226/INA826 | G = 1 + 49.4K/Rg | Rg = 249(1%)200倍
| 模数转换 | ADS11150 | 16位 工作电压2.7V-5.5V | 最大输入Vp-Vn 6V
|  | MCP3221 | 12位 工作电压2.7V-5.5V | 单位cN不符合要求
串口输出1 | 触摸屏modbus通讯
串口输出2 | 串口打印机
串口输出3 | modbusPC通讯
输入电源1 | 12V | 传感器10V
| 5V | IC
| 3.3V | IC/ARM
输入电源2 | 24V | 步进驱动

## 功能
试验参数设置
| 每批次数 | 3
| 试验批数 | 1
| 拉伸速度 | 500mm/min
| 试样细度 | 074.0tex
| 夹持长度 | 200mm
| 定长 | %
| 定力 | cN
显示内容 上半页
| 强力： | cN | 第一批第一次
| 伸长： | mm | 总共1批，每批3次
| 时间： | s |  |  | Strong 强力
| 显示内容 下半页 |  |  |  | Length 伸长
| 强力： | cN |  |  | Strength 强度
| 时间： | s |  |  | Elongation 伸长率
| 伸长： | mm
| 伸长率 | 0%
| 强度： | 0.00cN/tex
| 丝杠螺距 | 5MM | MOVING 长度校准 | 1、up方向到限位
| 驱动细分可调 |  |  | 2、启动down方向到指定长度
|  |  | 拉力测试 | 第一次
ES、收到测试命令，判断测试池是否有动作未完成，测试是否完成、未完成重复动作1
1、down向下拉伸
2、AD值突降，STOP立即停止
3、up到起始位置
力值校准
AD传感器校准
| 零度校准 | 记录0值时传感器值
| 满度校准 | 记录指定标称重力值
| 校正力度满度 | example 1kg砝码 9.8N | 2kg 19.6N
| ad | 显示单位设置
2kg | 1200 | N | 1960/19.6N
2kg |  | cN
2kg | 1200 | kg | 2000/20kg |  |  | MAX
2kg |  | lb
测试结果报表
语言设置

## ADS1115
| 测试系统 | VDD | PGA[2:0] | MODE | DR[2:0]: Data rate
|  | 3.3V | 010 : FS = ±2.048V (default) | 0 : Continuous conversion mode | 111 : 860SPS
| 批量系统 | VDD | PGA[2:0] | MODE
|  | 5V | 001 : FS = ±4.096V | 0 : Continuous conversion mode | 111 : 860SPS
For example, to write to the configuration register to set the ADS1113/4/5 to continuous conversion mode and then read the conversion result, send the following bytes in this order:
Write to Config register:
First byte: 0b10010000 (first 7-bit I 2 C address followed by a low read/write bit)
Second byte: 0b00000001 (points to Config register)
Third byte: 0b10000100 (MSB of the Config register to be written)
Fourth byte: 0b10000011 (LSB of the Config register to be written)
Write to Pointer register:
First byte: 0b10010000 (first 7-bit I 2 C address followed by a low read/write bit)
Second byte: 0b00000000 (points to Conversion register
Read Conversion register:
First byte: 0b10010001 (first 7-bit I 2 C address followed by a high read/write bit)
Second byte: the ADS1113/4/5 response with the MSB of the Conversion register
Third byte: the ADS1113/4/5 response with the LSB of the Conversion register

## H_bate1
I2C1_SCL | PB8
I2C2_SDA | PB9
TIM1_CH1 | PA8
TIM1_CH2 | PA9
TIM1_CH3 | PA10
TIM1_CH4 | PA11
TIM1_CH1N | PA7
TIM1_CH2N | PB0
TIM1_CH3N | PB1